DISC  95

Digital Interchange

Standards for Cardiology

A Demonstration of the DICOM Interchange Media

Sponsored by the American College of Cardiology

in Cooperation with the American College of Radiology and the

National Electrical Manufacturer s Association

Version 1.02

April 3, 1995

Software and documentation written by

Jonathan L. Elion MD, FACC

1994, 1995 Brown University Institute for Medical Computing and

the American College of Cardiology

Table of Contents

.Begin Table C.

Introduction	1

Creating Data Files for the Demo	2

Understanding X-Ray Angiographic IOD s	2

Other IOD s	4

Script File Philosophy	6

Script File Syntax	6

Generating XA IOD Files From Raw Pixel Data	8

Explanation of the UID s Used in the Test Data	9

Understanding Ultrasound IOD s	10

Accessing Files on the CD ROM s	11

Utilities	12

Build Program	12

Parse Program	15

The Demo Program	18

General Description	18

Globals	19

addToList()	20

badFile()	20

decompFrame()	20

displayImage()	21

displayRow()	22

initGUI()	22

initList()	22

main()	22

parseDicomDir()	23

parseDicomInfo()	23

parseDir()	23

readBytes()	24

readFileMetaInfo()	24

readString()	25

readTag()	25

selectFromList()	25

DICOM: Interpreting and Reviewing the Standard	26

Frequently Asked Questions (FAQ)	30

.End Table C.

Notice

The software and support documentation for the American College of Cardiology s 1995 demonstration of DICOM interchange media was developed at the Brown University Institute for Medical Computing.

This software is made available  as is  and neither the American College of Cardiology (ACC) nor Brown University make any warranty about the software, its performance, its merchantability or fitness for any particular use, freedom from defects, or its conformance to any specification.  The entire risk as to quality and performance of the software is with the user.

Copyright of the software and supporting documentation is jointly owned by the ACC and Brown University, and free access is hereby granted as a license to use this software, copy this software and prepare derivative works based upon this software.  However, any distribution of the source code or executable versions of this software, or of the supporting documentation (or of derivative works of the software or supporting documentation) must include this copyright notice and explanation.

Introduction.c.Introduction;

The American College of Cardiology (ACC) has been working with the National Electrical Manufacturer s Association (NEMA) and the American College of Radiology (ACR) to help develop standards for the storage and exchange of x-ray angiographic data.  The X-Ray Angiographic IOD Ad-Hoc Committee, composed of representatives of the above three groups, has been meeting to extend the DICOM 3.0 standard developed by ACR and NEMA to include the necessary specifications for x-ray angiography.  After considering a wide variety of currently available technologies, recordable Compact Disc (CD R) was selected as the physical medium to be used for exchanging images.

The first demonstration of  the exchange of x-ray angiographic and echocardiographic data among vendors was held in conjunction with the annual scientific sessions of the ACC in New Orleans in March, 1995.  By consensus of the groups mentioned above, the demonstration was restricted mainly to the  required fields (Type 1 and Type 2) for those modules of the single plan X-Ray Angiography (XA) and Ultrasound (US) that are mandatory.  The demonstration is known as  DISC 95" (Digital Interchange Standards for Cardiology).

The intent of this demonstration is to simulate how a CD R can be used for exchanging images on a single patient between institutions.  Therefore, images will be presented as if they came from a single patient (in reality, a variety of images from different vendors will be included, but listed in the directory as if they all came from the same patient).

Technically, the interchange medium selected is CD R, meaning that a single platter is recorded by a single writing device.  The disk that is produced is compatible by standard CD ROM readers.  The data can also be reproduced using the manufacturing techniques developed for traditional CD ROM s.  This demonstration will use a large-scale  pressing  of 5,000 CD ROM s of the file set so that adequate numbers of the discs are available for meeting attendees, vendors, and ACC officials.  The differences between a CD R and a CD ROM are technical with regards to their method of recording, and will not effect the reading of the discs, or the demonstration of support for the DICOM standard.

The software to support the demonstration has been written by the Brown University Institute for Medical Computing, and was supplied to participating vendors to assist with their preparation for the demonstration.  The software has now been placed in the public domain, with copyright maintained by the Brown University Institute for Medical Computing and the American College of Cardiology.  This process is intended to mirror that used by the Radiological Society of North America (RSNA) and Washington University s Mallinckrodt Institute of Radiology in the software to support the InfoRad demonstrations of DICOM networking.

Creating Data Files for the Demo.c.Creating Data Files for the Demo;

The software, data files, and support documentation are based around the DICOM 3.0 standard, and more specifically Supplement 4 on X-Ray Angiographic Image Objects and Media Storage, Frozen Version 1.5, from September 15, 1994.  The version of the DICOM standard used is labeled  STATUS: Final Text - October 29, 1993 .

Understanding X-Ray Angiographic IOD s.c.:Understanding X-Ray Angiographic IOD s;

This section of the documentation summarizes the process by which DICOM files of x-ray angiographic objects have been created.    The specific programs and data files that implement this process are described further in the sections that follow.  The reader is referred to DICOM Supplement 4 on X-Ray Angiographic Image Objects and Media Storage, and the general DICOM 3.0 standards for further details on the DICOM standard.  Appropriate sections of that document are referenced in the descriptions that follow.  While this section may appear at first to go into more detail than is necessary to understand the programs that comprise the demonstration, it is recommended that participants familiarize themselves fully  with the referenced sections of the standard, and with the concepts that are reviewed here.

The XA Image Information Object Definition (IOD) consists of a set of IOD modules that are included (see Table A.14.3-1 of the Part 3 Addendum on XA IOD Definition), some of which are central components of the existing DICOM standard, and some of which are extensions to accommodate x-ray angiography.  To collapse the XA IOD into a simple list of attributes, Type 1 and Type 2 attributes from the following mandatory modules have been included in the DICOM files that have been generated (Section refers to the section within Annex C of Part 3 that describes the Module, and Abbrev is an arbitrary abbreviation for the module):

Section	Abbrev	Description

C.7.1.1	PAT		Patient Module Attributes

C.7.3.1	SET		General Series Module

C.7.5.1	EQP		General Equipment Module

C.7.6.1	IMG		General Image Module

C.7.6.3	PIX		Image Pixel Module

C.7.6.4	CON		Contrast/Bolus Module

C.7.6.5	CIN		Cine Module

C.7.6.6	MFR		Multi─Frame Module

C.8.7.1	XAI		XA Image

C.8.7.7	XAP		XA Positioner

C.12.1		SOP		SOP Common

No information is included from the following mandatory modules, as the modules do not have any Type 1 or Type 2 fields:

Section	Description

C.7.2.1	General Study Module Attributes

C.7.6.10	Mask Module (no subtraction in the demo)

C.8.7.2	XA Acquisition

No information is included from the following modules, as their inclusion in the IOD is optional, and not required for this demonstration:

Module	Description

C.7.2.2	Patient Study Module

C.7.6.9	Frame Pointers Module

C.7.6.9	Mask Module

C.8.7.4	XA Table (required only if table moves)

C.8.7.5	XA Device and Therapy

C.9.2		Overlay Plane

C.9.3		Multi-Frame Overlay

C.10.2		Curve

C.11.2		VOI LUT

From these modules come the following list of elements that is specified for XA IOD s in this demo (this is a complete list, including a selected few Type 3 fields).

Element		Abbrev	Description

(0008,0008)	XAI		Image Type

(0008,0016)	SOP		SOP Class UID (XA IOD)

(0008,0018)	SOP		SOP Instance UID

(0008,0020)	STU		Study Date

(0008,0023)	IMG		Image Date

(0008,0030)	STU		Study Time

(0008,0033)	IMG		Image Time

(0008,0050)	STU		Accession Number

(0008,0060)	SER		Modality

(0008,0070)	EQP		Manufacturer

(0008,0080)	EQP		Institution Name

(0008,0081)	EQP		Institution Address

(0008,0090)	STU		Referring Physician's Name

(0008,1030)	STU		Study Description

(0008,103E)	SER		Series Description

(0010,0010)	PAT		Patient name

(0010,0020)	PAT		Patient ID

(0010,0030)	PAT		DOB

(0010,0040)	PAT		Sex

(0018,0010)	CON		Contrast/Bolus Agent

(0018,0060)	XAA		KvP

(0018,1063)	CIN		Frame Time

(0018,1152)	CIN		Exposure

(0018,1155)	XAA		Radiation Setting

(0018,1500)	XAP		Positioner Motion

(0018,1510)	XAP		Positioner Primary Angle

(0018,1511)	XAP		Positioner Secondary Angle

(0020,000D)	STU		Study Instance UID

(0020,000E)	SER		Series Instance UID

(0020,0010)	STU		Study ID

(0020,0011)	SER		Series Number

(0020,0013)	IMG		Image Number

(0020,0020) 	IMG		Patient Orientation

(0020,4000) 	IMG		Image Comments

(0028,0002)	PIX		Samples per pixel

(0028,0004)	PIX		Photometric Interpretation

(0028,0008)	MFR		Number of Frames

(0028,0009)	MFR		Frame Increment Pointer

(0028,0010)	PIX		Rows

(0028,0011)	PIX		Columns

(0028,0100)	PIX		Bits Allocated

(0028,0101)	PIX		Bits Stored

(0028,0102)	PIX		High Bit

(0028,0103)	PIX		Pixel Representation

(0028,1040)	XAI		Pixel Intensity Relationship

(7FE0,0010)	PIX		Pixel Data

Other IOD s.c.:Other IOD s;

In addition to the above, the following elements may be found in ultrasound IOD s on the demonstration disc:

Attribute		Description

(0008,2120)	Stage Name

(0008,2122)	Stage Number

(0008,2124)	Number of Stages

(0008,2128)	View Number

(0008,212A)	Number of Views in Stage

(0008,2204)	Transducer Orientation

(0018,1030)	Protocol Name

(0020,0060)	Laterality

(0028,0006)	Planar Configuration

(0028,0034)	Pixel Aspect Ratio

(0028,1111)	Large Red Palette CLUT Descriptor

(0028,1112)	Large Green Palette CLUT Descriptor

(0028,1113)	Large Blue Palette CLUT Descriptor

(0028,1211)	Large Red Palette CLUT Data

(0028,1212)	Large Green Palette CLUT Data

(0028,1213)	Large Blue Palette CLUT Data

(0028,1214)	Large Palette UID

The following elements are included in each IOD as part of the File Meta Information that helps to make up a DICOM file:

Attribute		Description

(0002,0000)	Group Length

(0002,0001)	Meta Information Version Number

(0002,0002)	SOP Class UID

(0002,0003)	SOP Instance UID

(0002,0010)	Transfer Syntax UID

(0002,0012)	Implementation Class UID

Finally, the following elements are included in the DICOMDIR:

Attribute		Description

(0004,1130)	File Set ID

(0004,1141)	File ID of File Set Descriptor

(0004,1142)	Format of File Set Descriptor File

(0004,1200)	Offset of First Directory Record

(0004,1202)	Offset of Last Directory Record

(0004,1212)	File-Set Consistency Flag

(0004,1220)	Directory Record Sequence

(0004,1400)	Offset of next directory record

(0004,1410)	Record in use flag

(0004,1420)	Offset of referenced lower-level directory entity

(0004,1430)	Directory Record Type

(0004,1500)	Referenced File ID

(0004,1510)	Referenced SOP Class UID in file

(0004,1511)	Referenced SOP Instance UID in file

(0004,1512)         Referenced Transfer Syntax UID in file

Script File Philosophy.c.:Script File Philosophy;

A simple utility (build) is used to translate ASCII  script  files containing DICOM attributes into the corresponding DICOM files.  In a sense, this is roughly analogous to the process of converting a program written in human-readable form into its corresponding binary representation.  It is not intended for the build program to have a significant understanding of IOD s, or even of DICOM protocols.  This knowledge is encapsulated in the script files.  No checking is done, for example, to see if a Type 1 field has been omitted.  It is the responsibility of the script-writer to assure the completeness of the data that is presented to the build utility.

There are two very strong advantages of this approach: simplicity and self-documentation.  The simplicity of the approach should be obvious.  The complexity of the build program is minimized, and a simple linear translation of ASCII to DICOM is accomplished.  The self-documentation aspect refers to the fact that the script file very completely describes (in human-readable form) the contents of the corresponding DICOM file.  Even a casual reader will immediately see what has been included in each file.

Since the build utility provides a simple linear translation, the information in the script files must be listed in ascending order by tag number (as described in Part 5, section 7.1 on Data Elements).  As a result of this requirement, information in the script files is no longer grouped by modules.  In the script files that have been created for the demonstration, comments have been added to each data item to document which module caused it to be included in the DICOM file.

Script File Syntax.c.:Script File Syntax;

Script input is read a single line at a time (input is line-oriented like FORTRAN or assembly language, and not stream-oriented like Pascal or C).  Blank lines are allowed, and lines may have unlimited white space (spaces and tabs).  Anything after a semi-colon is ignored (used for comments).  No spaces are allowed within fields (such as within a DICOM tag).  All characters on the line are converted to upper case.  The exceptions are strings, enclosed in "quotation marks", and may contain spaces and mixed upper and lower case (case is preserved).

Almost all lines of the input file will contain an attribute tag and a data value.  All attribute tags are considered to be hexadecimal constants, for example: (7FE0,0010).  It is not necessary to specify hexadecimal notation (0x7FE0,0x0010), as this is always assumed for tags.

Locations in the file can be referred to symbolically using labels.  A label appears by itself on a line, and is an alphanumeric string (no blanks) terminated with a colon, for example:

FIRST_PATIENT:

This symbolic addressing is helpful for specifying offsets (a capability essential for creating the entries in the DICOMDIR).  To reference an offset marked by a label in the data field, a  &  is used before the label name (for example: &FIRST_PATIENT).  Here is an example of the symbolic addressing capability to specify the offset of the first record of the root directory entity (offset are from the beginning of the preamble of the File Meta Information).  This example is based on DICOMDIR.SCP:

(0004,1130) "DISC95_DEMO"		; File─set ID

(0004,1200) &FIRST_PATIENT	; Offset of First Record of Root Dir

(0004,1202) &LAST_PATIENT		; Offset of Last Record of Root Dir

(0004,1212) 0				; File─set Consistency Flag

(0004,1220) 0xFFFFFFFF		; Dir Record Sequence (start of set)

FIRST_PATIENT:

(FFFE,E000) 0xFFFFFFFF		; Item Data Element

(0004,1400) 0				; Offset of next Dir Record

(0004,1410) 0xFFFF			; Record in use flag

It is also necessary to have a way to specify the length of a span of bytes (for example, the Group Length attribute of the File Meta Information).  In this case, a  #  is used to begin the symbolic reference, and two labels are given, separated by a colon.  This example is also taken from DICOMDIR.SCP, and shows the specification for group length:

(0002,0000)	#GROUP_START:GROUP_END ;	Group Length

GROUP_START:

(0002,0001) 0x0001			; File Meta Info version number

(0002,0002) 1.2.840.10008.1.3.10	; SOP Class UID (DICOMDIR)

(0002,0003) 999.999.2		; SOP Instance UID (dummy)

(0002,0010) 1.2.840.10008.1.2 	; Xfer Syntax UID

(0002,0012) 999.999			; Implementation Class UID

GROUP_END:

The BEGIN command opens the file for output and writes File Meta Information.  The name of the output file (used to write the binary version of the XA IOD) is the only other argument allowed on the line.  For example:

BEGIN XA000001

would open the file XA000001 to receive the binary version of the XA IOD generated.

The specification for a single XA IOD is terminated with an END statement which has no arguments.  For the demonstration data that is provided, only one XA IOD is specified per script input file.  It is possible to batch these together, so that a single input file contains specifications for several IOD files, each contained between its corresponding  BEGIN and END statement.

Generating XA IOD Files From Raw Pixel Data.c.:Generating XA IOD Files From Raw Pixel Data;

While it is not intended that the software provided to support DISC 95 function as a general-purpose DICOM tool, the build utility (described below) can be used to generate XA IOD files from raw pixel data.  This section describes some of the procedures used to generate the data files included in the distribution, in case additional test data is to be generated locally.  The descriptions in this section are restricted to tools used in the Unix environment.  Equivalent tools are often available in other environments, and in any case, these descriptions serve to document the procedures used.

The general steps needed to create the data files include:

Decompress and/or translate proprietary file formats, storing each sequence in a separate file (this step is not described further, as it is dependent on the specific digital archive or transfer system used to obtain images from the acquisition system).

Strip off the header (if any) from the sequence file, leaving only raw pixels.

Perform JPEG compression on the raw pixel file

Create an icon image

The following commands are taken from a Unix shell script that performs these tasks.  Several of the commands are based on the Utah Raster Toolkit (urt), which can be obtained from any of a variety of anonymous FTP sites, such as ftp.uu.net in subdirectory graphics/urt, or at cs.utah.edu, in subdirectory pub.

The image sequence starts in file cine001.digital.  First, tail is used to skip the 2K header (have to start on byte 2049), and the results (raw 512 x 512 pixel frames) are put into cine001.pix.  Then, the raw frames are JPEG compressed, using the seq2ljpg utility with a PSV value of 1 (-S1).  The -t option generates an offset table, with the compressed images going to cine001.jpg, and the offset table going to cine001.tab.  As a double-check, the frames are decompressed to be sure the sequence file is the same as we started with.  If files are identical, the cmp command will print no message.

#!/bin/csh

tail +2049c cine001.digital > cine001.pix

seq2ljpg ─S1 ─t cine001.pix cine001.jpg cine001.tab

ljpg2seq cine001.jpg cine001.compare

echo "File comparison:" `cmp cine001.pix cine001.compare`

rm cine001.compare

Now pull out frame 55 for an icon image.  The dd command works on a default block size of 512 bytes.  Therefore, each frame is 512 blocks, and the 55th frame starts on block 28160.

dd if=cine001.pix of=cine001.055 skip=28160 count=512

The next pipeline converts the image to run-length encoded (the native format used by the Utah Raster Toolkit) with one channel (grayscale).  The image then gets flipped, scaled to 128 x 128, and converted to a Sun raster file (easiest way to get a single-channel output file).  The fant command that does the scaling gets things one pixel too big when scaled 25% (-s 0.25); the 0.247 factor seems to work well, and yields a proper 128 x 128 icon image.  The xloadimage utility displays the icon image on a X-Windows display, just to help verify that it was a good frame to select.  Finally, the tail command strips off the Sun raster header and we are left with a raw 128 x 128 icon file.

rawtorle cine001.055 ─w 512 ─h 512 ─n 1 rleflip ─v \

| fant ─s 0.247 0.247 | rletorast ─o cine001.rast

xloadimage ─zoom 400 ─brighten 150 cine001.rast

tail +33c cine001.rast > cine001.icn

Explanation of the UID s Used in the Test Data.c.:Explanation of the UID s Used in the Test Data;

The scripts files that are used to generated the testing versions of the XA IOD s contain several UID s, some of which are defined by the DICOM standard, and some of which are locally generated.  The latter are intended to simulate the UID s that might be used by a typical institution in DICOM activities.

The following official DICOM UID s are used:

1.2.840.10008.1.3.10			SOP Class for DICOMDIR

1.2.840.10008.5.1.4.1.1.12.1	SOP Class for XA IOD

1.2.840.10008.5.1.4.1.1.3		SOP Class for Multi-Frame Ultrasound IOD

1.2.840.10008.5.1.4.1.1.6		SOP Class for Ultrasound IOD

1.2.840.10008.1.2.1			Transfer syntax for Little Endian Explicit

1.2.840.10008.1.2.4.70		Transfer syntax for Lossless JPEG 14

The fictitious hospital that generated the studies and the files for the CD R is identified by the institutional UID of 999.999.  This is not a valid (registered) UID, and is used for demonstration purposes only.  It would be typical for an institution (or vendor) to develop an algorithm that generates unique UID s for studies, series, images, etc.  A simple algorithm is presented here that is not intended to reflect that of any particular real-world installation, nor is it necessarily recommended for implementation.  In fact, the algorithm does not totally guarantee that UID s would be unique, as it is possible (although unlikely) that two studies generated at exactly the same minute on different devices could both have the same UID number generated.  We will assume that the hospital in the demonstration only has one instrument generating DICOM files, and the algorithm therefore adequately assures uniqueness.

The organization identified by UID 999.999 will use the prefix 999.999.1 for XA IOD instances, 999.999.2 for ultrasound, and 999.999.2 for  DICOMDIR files.  XA Study Instance UID s are created by appending the study date and study time (leading zeroes must be suppressed).  Series UID s are created by appending the series number to the Study Instance UID, and the image instance UID will be created by further appending the image number.

Understanding Ultrasound IOD s.c.:Understanding Ultrasound IOD s;

While the major thrust of the development effort for this demonstration has been geared at X-Ray Angiography, there are ultrasound images that are being included on the DISC 95 demonstration disc.  This section provides a very brief discussion of some of the nuances of dealing with these files.  Most of the methods for handling and displaying ultrasound information from the DICOMDIR or from the IOD files themselves are very similar to the methods used for XA IOD s.  The major difference is in the way that the pixel data is stored.  For this demonstration, only the following pixel formats are supported:

MONOCHROME2:  This format is nearly identical to the XA IOD s, except that the pixel data is not compressed, and the image size may not be a simple 512 by 512.

RGB:  The format is similar to MONOCHROME2, except that each pixel is represented by three bytes, not one.  These bytes represent the red, green, and blue components of the displayed pixel respectively.  All RGB images on the disc are stored in this format.  There is an alternative pixel storage scheme possible, known as  planar , where first all the red pixels are stored for a frame, then all the green pixels, and finally all the blue pixels.  This format is not used in the files on the disc.

PALETTE COLOR:  This approach (sometimes called  indexed color ) uses 8-bit pixels, but also has lookup tables that convert the 8-bit values into red-green-blue three-byte values.  Once the 8-bit pixel value is converted in this fashion, it can be handled much like the RGB pixels.

Accessing Files on the CD ROM s.c.:Accessing Files on the CD ROM s;

Every attempt has been made to assure that the directory and file formats on the CD ROM match the ISO 9660 and DICOM standards.  There are a few  quirks  of file system interoperability that have been encountered, and are described here.

An ISO 9660 Volume Descriptor is contained on the disc (but would not be available when using standard device drivers on Macintosh, DOS, Windows, or Unix).  One of the fields in this Descriptor is the name of a file containing copyright information.  The file COPYRITE. is named in this field, and included on the disc (note that DICOMDIR does not know of this file, and it is therefore technically not visible to a DICOM File Set Reader).  The spelling of the file is intentional, to assure strict adherence to the ISO 9660 file naming convention.

ISO 9660 version numbers are specified in the ISO directory, but get filtered out by traditional device drivers.  Furthermore, the device drivers on different operating systems will attempt to convert file names encountered into something well understood by the target operating system.  These leads to some unexpected results, as described below (along with other operating system dependencies).

DOS and Windows: The upper and lower case content of file names is not relevant.  The use of the  .  when there are no file types is optional.  Therefore,  DICOMDIR ,  DICOMDIR.   dicomdir  and  dicomdir.  are all equivalent.  The README file on the disc is pointed to by DICOMDIR, and is visible to File Set Reader applications.  As specified in the DICOM standard, its lines are terminated by CR/LF pairs (very DOS-like!).

Macintosh: The standard ISO file name conversion routines show the file names as all upper case, and with the final  .  stripped off.  The name of the volume of the disc is DISC95_DEMO (visible on the desktop when the disc is mounted).  The presence of the CR/LF pairs in the README file will cause a problem when a program like TeachText is used to display the file, as the LF s show up as a   on the screen.  The LF s should be filtered out before display.

Unix: File names are converted to all lower case, and the final  .  is left intact.  The presence of the CR/LF pairs in the README file causes a problem in that the CR s are visible (often showing up as ^M, depending on the program used for display).  These must be filtered out before display.

Utilities.c.Utilities;

This section of the documentation is currently restricted to a description of the requirements to run each of the utility programs supplied.  Technical descriptions of the internal workings of each program is provided through the extensive comments within the source code of each program.

Build Program.c.:Build Program;

The build program scans a script file (described above) and produces the corresponding XA IOD file.  The program is invoked as:

build <script file>

where <script file> is the name of the input script file.  There are no command line options.  The output file (or files) are created based on specifications within the script file, and not anything on the command line.  The files distributed with the demo software use one script file for each XA IOD file created, but theoretically these could all be combined into a single script file.

Most error messages are generated during the scanning of the script file, and are of the form:

Error on line nn of script file: msg

where nn refers to the line number of the input script file, and msg is a brief explanation of the problem encountered.  The following is a more complete list of the error messages that may be generated, along with their more common causes, and the corrective action that is required.

'*' only valid in (7FE0,0010)

The  *  specification refers to a powerful feature for including encoded pixel files into the XA IOD.  It may only be used with the (7FE0,0010) attribute.

(gggg,eeee) not in dictionary

The dictionary module contains what  should be a complete list of the group and element tags that are needed, at least for this demo.  The group gggg and element eeee listed could not be found in the dictionary.  Either a typographical error was made in entering the data, or a new element has been added to the script that has not been defined in the dictionary.  In the latter case (which should not be occurring for this demonstration) add the corresponding entry to the dictionary source code and recompile the build utility.

BEGIN expected

Valid data is contained in a BEGIN block, started with the BEGIN statement and ending with the END statement.  The scan did not find a BEGIN statement where expected.  Possibly the BEGIN statement is misspelled, or missing altogether.

Duplicate label

The same label has appeared twice with a BEGIN block.  Pick another name for this  location in the file.  Labels can be referenced any number of times, but only defined once.  The names are local within a BEGIN block.

Error closing auxiliary input file

An error condition was raised by the fclose() of an auxiliary input file. This may occur during processing of a file specification in OB attribute.

Error closing offset table file

An error condition was raised by the fclose() of an offset table file. This may occur during processing of a file specification in OB attribute.

Error closing output file

An error condition was raised by the fclose() for the output file (the name of the output file is specified on the BEGIN statement).

Improperly formed hex constant

A hex constant was expected, but a leading 0x was not found.

Improperly formed tag

A tag must be expressed as (gggg,eeee), where gggg and eeee are 4-digit hexadecimal numbers, and there is no white space.  A common error is to leave a space after the comma.  The error may also be generated if gggg or eeee is not a valid hex number (0-9 and a-f).

No script file specified

The name of the script file must be given on the command line.  This message indicates that no corresponding command line argument was found.

Offset table too small: change #define MAX_OFFSETS

In the demonstration version of the software, the offset table is a fixed size, and will accommodate multi-frame image files with up to 500 frames.  This value must be increased at compile time for the utility, by changing the #define MAX_OFFSETS 500 line in acc95.h to a higher value.

Too many fields on the line

An input line of the script generally consists of only the tag and corresponding value (this would be two fields on a line).  This message can be generated by an improperly formed input line, such as a text string not surrounded by quotes (a string with quotation marks counts as a single field) or a missing semicolon before a comment.

Too many symbols: change #define MAX_SYMBOLS

In the demonstration software, the maximum size of the symbol table is determined at compile time, as it is handled as an array.  If more than 100 symbols are required, the value must be increased a compile time for the utility, by changing the #define MAX_SYMBOLS 100 line in acc95.h to a higher value.

Unable to open auxiliary file fff

An error has occurred while trying to open the auxiliary file fff for input.  This may occur during processing of a file specification in OB attribute.

Unable to open output file fff

An error has occurred while trying to open file fff for output.  This may be seen during the processing of a BEGIN statement if an I/O error occurs while trying to open the file for output (for example, no write permission for the directory).

Unable to open offset table file

An error occurred while trying to open the offset table for an encoded pixel specification.  This may occur during processing of an OB attribute.

Unable to open pixel file

An error occurred while trying to open the pixel file for an encoded pixel specification.  This may occur during processing of an OB attribute.

Unable to open script file fff

An error occurred while trying to open script file fff for input.  Either the read permissions on the file need to be changed, or the file name has been misspelled.

Undefined symbol

A reference has been made to a label that does not appear in the source file (actually within BEGIN/END blocks).  Check your spelling, or add the necessary label.

Uneven byte count in hex constant

Hexadecimal constants can be specified in several places (as offsets, as values to be included in OB attributes, etc.).  Hexadecimal constants must always be specified with an even number of bytes.

Unexpected End─Of─File

End of file was reached on the script file while looking for more lines to scan.  Most likely due to a missing END statement.

Unknown VR type

When sending DICOM data, implicit Value Representation (VR) mode is used.  The actual VR is looked up in the dictionary table, and used to guide the formatting of the data to the output stream.  This message should not occur for this demonstration.  It indicates that another case in the switch statement has to be added for the VR.

Unmatched quote

Text strings are demarcated by a quotation mark at the beginning and end of the string.  Text cannot carry over from one line to the other.  This message indicates that a text string was found (an initial quotation mark), but that there was no corresponding closing quotation mark.

Parse Program.c.:Parse Program;

This is a simple program that takes an XA IOD (or most any DICOM file) and prints its contents (almost like a  dis-assembler ).  The program is invoked as:

parse <DICOM file> [<output file>]

where <DICOM file> is the name of the DICOM file to be parsed and formatted, and  <output file> is an optional file name to be used for writing pixel data.  There are no command line options.  Here is an excerpt from the output of the program as run on the XA001 file of the demo data:

00000182:  (0008,0020) [DA/8]	"19940822"	; Study Date

00000192:  (0008,0030) [TM/6]	"083000"		; Study Time

000001A0:  (0008,0050) [SH/0]	""			; Accession Number

000001A8:  (0008,0060) [CS/2]	"XA"		; Modality

000001B2:  (0008,0070) [LO/14]	"Acme Products"	; Manufacturer

000001C8:  (0008,0090) [PN/14]	"Localdoc^Joan"	; Referring Physician's Name

000001DE:  (0010,0010) [PN/14]	"Brown^Charles"	; Patient's Name

000001F4:  (0010,0020) [LO/12]	"123─45─6789"	; Patient ID

00000208:  (0010,0030) [DA/8]	"19940822"	; Patient's Birth Date

00000218:  (0010,0040) [CS/2]	"M"			; Patient's Sex

A very handy feature has been added to the program to allow DICOM files to be parsed, then edited and re-built.  When the option file name is given on the command line, this is used as the basename for building file names for output.  A DICOM file with a single pixel stream in it would have its pixels written to <filename>.000 (and the source listing printed on stdout).  Other files can be generated during the parsing process, for example, palette LUT s in a PALETTE COLOR image.  In this case, the binary data files are written in the order in which they are encountered, so that the LUT s end up in <filename>.000, <filename>.001, and <filename>.002, while the pixel data ends up in <filename>.003.  Another side-effect of providing a file name on the command line is that the format of the output will be changed slightly to conform to the input requirements of the build utility.

Here is an example of how this feature can be used to look at a DICOM file unknown.dcm:

parse unknown.dcm xa000083 > xa000083.scp

build xa000083.scp

In this example, the pixels from the unknown.dcm file are put in xa000083.000.  The output of the parse utility is redirected into xa000083.scp, which could then be edited to make any necessary changes (or examine the contents of the file).  When the build utility is simply run immediately on the unedited version of xa000083.scp, the resulting XA000083 DICOM file should be identical to the starting file unknown.dcm.  This capability of the parse utility was used during the preparation of the DISC 95 demonstration disc to edit DICOM files sent from a variety of vendors

There are several error conditions that can be detected.  The following is a list of error message that may be reported, along with a more complete description of their cause and/or corrective actions:

Can't handle VR type: ttt

This error should not occur, as it indicates that a VR type was encountered in the input file that does not have a matching case in the switch statement that formats the various types of values.  The VR type found in the input file was ttt.

Error: Not a valid DICOM file.  Will try to read as non─Part 10 file.

A DICOM file begins with a 128-byte preamble (that is ignored by this program), followed by the 4-byte character string DICM.  In this case, this is not what was found.  The input file stream is rewound, and an attempt is made to read and format the input without looking for the preamble.  This can occur if a so-called DICOM file was created that does not include the File Meta Information as specified by Part 10.

Error closing aux output file.

An unexpected error was encountered when closing one of the output files (when the optional file name is given in the command line.

No input file specified

The name of the DICOM file to be formatted must be included as the sole command-line option.

Unable to open input file fff

The input file fff could not be found or could not be opened.  Either the file does not exist, its name has been misspelled on the command line, or the user does not have read access to the file.

Undefined tag: (gggg,eeee)

An element was encountered in the file that could not be located in the program s dictionary (contained in dict.c).  This can happen when invalid data is in the file being parsed, or if the dictionary is incomplete (for example, as DICOM application profiles are expanded).

Unexpected EOF

An end-of-file condition was encountered in the middle of reading information from the file (at any time other than trying to read the next tag).  This may indicate invalid data in the DICOM file, or a problem with the parse program.

Warning: Explicit VR tag ttt does not match dictionary VR ttt

An attribute has been read with Explicit VR, but the VR found in the DICOM file does not match the VR associated with the tag in parse s dictionary.  This is most likely caused by an error in the program that generated the DICOM file, as the dictionary used by parse was taken directly from the Part 6 data dictionary file of the published standard (it was extracted from the machine-readable version of the standard).  The first ttt tag is the one encountered in the DICOM file, the second is the expected value based on the dictionary.

The Demo Program.c.The Demo Program;

The demo program demo.c is intended to be the starting point for a developer s program to display images from the CD R for the DISC 95 DICOM demo.  The main purpose of the program is to show how to interact with the information contained in the DICOMDIR file.

The program does very little error checking on the content of the DICOMDIR or the IOD files, as it is assumed that reasonable DICOM files are present.  This is true for the DISC  95 disc, as the data has been thoroughly examined and verified by many groups prior to the pressing of the CD ROMs.  Obviously, considerable error checking and handling should be added if this is to be turned into commercial-grade product code.  One thing in particular that is not checked is the order of the DICOM tags (by definition, they must be presented in ascending order).

This program ignores many of the fields in DICOMDIR, but does extract and store the information (rather than just skip over it).	 This is intended to make it easier to modify the program to use the information if you are interested in it.

It is hoped that significant portions of this code will be incorporated into demonstrations or into commercial products.  In that case, it is important to remember to preserve the presence of the original copyright notice, with developers adding their own.  Standard practice is for the copyright notice to be printed on the screen when the program is first invoked.  It should also appear in the source code and in the resulting binary files.

To modify this code, you will have to (at the very least) change the code that provides the user interaction for the selection process (presumably, you will want this to use the Graphical User Interface (GUI) of your target display system).  You will also have to write code that moves the image pixels to the display s frame buffer, and also code to handle of the icon images (if you choose to implement this feature).

The program has been written in a highly modular fashion to facilitate these changes.  In many cases, code is contained in subroutines even though the subroutine is only called once from a single place in the program.  This is done to facilitate readability (keeping main program modules small), and to isolate and identify areas that may require modification for specific target systems.  The version of the demo.c program that is distributed contains simple ASCII-based versions of some of the routines (for example, selection of an image from a list).  This should allow an understanding of the structure of the program while avoiding platform-specific code.

General Description.c.:General Description;

The information contained within DICOMDIR is formatted as a multiply-linked list that implements a directory tree.  There is an initial pointer to the first directory record (the first patient).  You could then follow the directory chain at the patient level, or follow a pointer down to the next level (study).  each level has similar links; one to the next directory record at the same level, and one to the next level down.

This program is designed to  flatten  the directory hierarchy into a simple linked list.  Each record in this list uniquely describes all of the directory information that has been provided about a given image file(remember that the DICOMDIR contains a number of fields of descriptive information at various levels.

Each flattened record in this linked list is described by struct dirRecord (all the structures mentioned here are defined in demo.h).  Within the dirRecord  structure, the information from the various levels (patient, study, series, image) is grouped together by the structures patientInfo, studyInfo, seriesInfo, and imageInfo respectively.  This grouping simplifies the task of zeroing out a lower level s information when the tree traversal routine comes back up a level. The list of these records is pointed to by struct dirRecord *dirList.  A single work area (struct dirRecord thisRec) is kept for constructing a flattened record.

Globals.c.:Globals;

Use of global variables has been kept to a minimum (but not totally eliminated as some programming styles would encourage).  These few globals are all accessed within this single source file.

struct dirRecord *dirList = (struct dirRecord *) NULL

A pointer to the start of the linked list of flattened directory records.

struct dirRecord thisRec

A work area for building up the flattened record.  This is constructed as the directory tree is traversed.

int choiceCount

A counter used by list selection routines.

FILE *inFile

The current input file (the DICOMDIR file or an XA IOD file)

char filesetID[VR_CS_LENGTH + 1]

The name of the DICOM file set.  This is the only piece of descriptive information kept from the introductory section of DICOMDIR.  The file set refers to the collection of DICOM files and their associated DICOMDIR.

int thisEndian

Endianism of the host CPU (BIG_ENDIAN or LITTLE_ENDIAN)

int explicit_VR

Does the transfer syntax of the input file use Explicit VR or Implicit VR

int encodedPixels

Does the transfer syntax of the input file use encoded pixels (compression)

int rowNumber

Row number in the image that is currently being decompressed

addToList().c.:addToList();

The addToList() routine is called repeatedly, once for each image found on the media.  Typically, the data from the dirRecord structure would be formatted into a line suitable for inclusion in the scrolling selection box.  The approach implemented in this demo gives a  flattened  view of the media, namely, there is a single list of all of the images on the media.  Alternatively, you might want a more structured approach (pick a patient, then pick the study, then series, then image).  A simple example of this type of hierarchical selection scheme is shown in the demo1.c program.  It is useful to note that the record being added to the selection at this point contains an icon image (128 x 128) that may be helpful in constructing the user interface for the selection.  This icon image is pointed to by recordToAdd─>imageInfo.icon[128][128].  The routine is declared as:

void addToList(struct dirRecord *recordToAdd)

badFile().c.:badFile();

The badFile() routine is used to bail out from reading the DICOMDIR file when an error is found.  An error message is printed, and the program exits.  Simple!  The routine is declared as:

void badFile(char *msg)

decompFrame().c.:decompFrame();

The decompFrame() routine is contained in jpeg.lib (in the jpeg subdicrectory), and is used to perform lossless JPEG decompression.  This code is a modified version of the public JPEG decompression software from Cornell (which in turn is modified from the code from the Independent JPEG Group).  This routine reads directly from the FILE *inFile of the demo program, which must be left positioned at the JPEG file header before decompFrame() is called.  The routine reads the JPEG data, parses the file and scan header, and decompresses the image data.  As each row of pixels becomes available, the displayRow() routine is called, passing it to the row to display.  The routine is contained in the decomp.c source file in the jpeg directory, and is declared as:

void decompFrame(void)

displayImage().c.:displayImage();

The displayImage() routine is used to display a multi─frame image.  The version that is included is somewhat of a  bare shell  of a routine, as the types of things that might be done with the image is highly vendor─specific.  This routine shows how information can be parsed from the XA IOD file, and how pixel data can be accessed.  The SOP Class UID in the File Meta Information is examined to see if the file is an XA IOD.  If not, it is ignored.  Attributes from the XA IOD are parsed, and a simple set of if statements allows examining and storing values from those fields considered interesting for display.  For the demo version of the software, only a few of the attributes are checked.  For product─level code, a more comprehensive set of attributes will need to be handled.

There are several ways that these pixels might have been written into the IOD:

1.	The (7FE0,0010) tag has a length that is NOT 0xFFFFFFFF.  This should not happen for an XA IOD, as it means that the pixel data is not encoded (the standard states that XA IOD s will use lossless JPEG).   But, for the sake of argument, this might happen during the testing phase of display software (if that's how you built your XA IOD file).  Frames can be read by reading 512*512 bytes at a time (one frame's worth).

2.	The (7FE0,0010) tag has a length that is 0xFFFFFFFF.  The next attribute should be (FFFE,E000), a basic offset table.  It can have a length field of 0x00000000 (meaning no offset table information is given), or may contain concatenated 32─bit  unsigned integer values that are byte offsets to the first byte of the Item Tag of the first fragment for each frame in the Sequence of Items. (See Part 5, Section A.4).  For this version of the software, the offset table is not included (length of 0x00000000).

3.	Given the situation in #2 above, the pixel data can be included in one or more fragments.  As of the beta release of the software, the encoded pixel data is divided up so that each frame is written to a separate fragment.  A sample implementation of JPEG decompression is provided, calling the decompFrame() routine that is derived from the public domain version of the lossless JPEG decompression.

The routine is passed the flattened record of the image file to be displayed, and is declared as:

void displayImage(struct dirRecord *oneRecord);

displayRow().c.:displayRow();

The displayRow() routine is called by the JPEG decompression routines to display a row of pixels.  The decompression routines work a row at a time, and calls this routine as each row is available for display.  The routine is passed the row of pixels to display, and  is declared as:

void displayRow(unsigned char *row)

initGUI().c.:initGUI();

The initGUI() routine initializes the Graphical User Interface environment.  This is called at the beginning of the execution of the program to set up display windows, and other routine chores that are part of setting up the GUI environment.  When modifying the code, be sure to include proper copyright credits (and add your own, of course).  The routine is declared as:

void initGUI(void)

initList().c.:initList();

The initList() routine is called to set up a list of choices for the user selection process.  Typically, this would be used to set up a selection window, initialize a pointer to the head of the selection list, etc.  It is possible that for a given GUI environment, no specific action needs to be taken.  The routine is declared as:

void initList(void)

main().c.:main();

The main() part of the program is written as  bare bones  code that is intended to show how the subroutines are used to manipulate the data in the DICOM file set.  It does not examine command line parameters.  This version of the control loop allows the user to repeatedly make selections from the list that has been constructed.  For each selection made, a call is made to the routines to allow display of the image.  If during the selection routine, the user requests to exit the program the selection routine returns NULL.  For a demonstration at a show, this option would probably not be offered (don t want to allow the demo software to be exited so easily).  The routine is declared as:

int main()

parseDicomDir().c.:parseDicomDir();

The parseDicomDir() is a simple control loop that traverses the DICOMDIR directory tree creating a flattened record.  The nested while() loop implements a depth─first search of the directory tree, extracting first patient information, then study information, series information, and image information.  The routine is passed the offset of the first patient record, and is declared as:

void parseDicomDir(UnsignedLong firstPatientPtr)

parseDicomInfo().c.:parseDicomInfo();

The parseDicomInfo() routine reads the introductory information from DICOMDIR.  The file consistency flag is checked to be sure it is 0 (otherwise an error is reported).  The only information currently kept for the informational section of the DICOMDIR is the name of the file set.    A pointer to the first directory record (patient record) is returned.  This is actually an offset value representing the location of the first patient record relative to the beginning of the DICOMDIR file.  The routine is declared as:

unsigned long parseDicomInfo(void)

parseDir().c.:parseDir();

The parseDir() routine is called to extract information from a single DICOMDIR directory record and save it into the flattened record (struct dirRecord thisRec).  The routine returns the offset of the next directory record at this level.  It is passed the offset of the directory record to be parsed, and the value of the record type that is expected (PATIENT, STUDY, SERIES, or IMAGE).  If it is an image record that is being parsed, then the completed flattened record add it to the linked list.  The routine also has the side-effect of setting the offset to the next level down (this is passed as a pointer to the variable to be set with this value).

This is another example of code that handles known and unknown lengths.  If a  valid length was given in the data item element, there will be no final item tag, and the length can be used as a guide to how much data to read.  If the length was 0xFFFFFFFF, then the length is unknown, and a final item tag marks the end of useful data.  So, this loop will continue until either the length of data remaining to be read is decremented to 0, or until a terminating tag is found.

The routine is declared as:

unsigned long parseDir(

unsigned long thisDir,		/* Dir record to be scanned */

unsigned long *lowerDir,		/* Return offset of next level*/

char *recordType)			/* Type of record to scan */

readBytes().c.:readBytes();

The readBytes() routine is used to read bytes from DICOMDIR.  It is passed a pointer to the destination for the data, the length of the data field to read, and a flag indicating whether an adjustment for the Endian status of the machine needs to be made.  If this flag is set to NO_ENDIAN, or if the code is run on a Little Endian machine, then the bytes are read indirectly.  If, however, the flag is set ENDIAN and this is not a Little Endian machine, then the byte order is reversed.

The routine is declared as:

int readBytes(char *value, long length, int flag)

readFileMetaInfo().c.:readFileMetaInfo();

The readFileMetaInfo() routine reads the DICOM File Meta Information from the input file.  The content of the File Meta Information is defined in Part 10, table 7.1.1.  After the 128─byte preamble and the 4─byte prefix are read, the remaining portion of the File Meta Information Header is read assuming the Implicit VR Little Endian Transfer Syntax (defined in Part 5).  Even the DICOMDIR is a DICOM file, and as such, must have its own File Meta Information header.

This routine uses a reading technique that is similar to one used for reading data items whose length may or may not be known.  In this case, the Group Length (a required field) is used as an indicator of how much data to read.  This length value is set initially at 9999, but quickly gets set to the correct length (the length recorded in the File Meta Information segment is the length of all Meta Information after the Group Length).  This little programming short-cut greatly simplifies the control loop used to read this information.

The routine is passed a pointer to a structure to hold the parsed information, and is declared as:

void readFileMetaInfo(struct metaInfoStruct *metaInfo)

readString().c.:readString();

This is a simple subroutine to read strings from a DICOM file.  Strings must be padded to be an even length (the padding is with a space 0x20 except for UID s, that are padded with a 0x00).  If there is a trailing 0x20 detected after the string is read, it is removed.  In any case, the string in the receiving buffer is terminated with a 0x00 (in case the calling routine did not zero out the buffer before calling for a string to be read).  The routine is passed a pointer to the location to receive the string, and the length of the string to be read, and is declared as:

void readFileMetaInfo(char *value, long length)

readTag().c.:readTag();

The readTag() routine is simply a short─hand for reading the group, element, and length fields from a DICOM file.  It is passed pointers to the variables to be set to the group, element, and length values.  It is declared as:

void readTag(unsigned short *group, unsigned short *element,

UsignedLong *length)

selectFromList().c.:selectFromList();

The selectFromList() routine should allow the user to interact with the list and make a selection.  The routine should return a pointer to the directory record selected (or NULL if the user wants to exit the program).  At a demonstration at a show, you probably do not want to allow the user to exit (at least, no so easily!).  The routine is declared as:

struct dirRrecord *selectFromList(void)

DICOM: Interpreting and Reviewing the Standard.c.DICOM: Interpreting and Reviewing the Standard;

In reviewing the DICOM standard and writing the software to support this demonstration project, several issues have surfaced regarding a few of the specifics of the standard.  No matter how carefully a standards body tries to anticipate all situations, it is inevitable that in the process of implementing software some uncertainties will arise.  It then is left to the implementor to interpret the intent of the standard.  This section of the documentation describes some of the issues that were encountered, and the approach taken in writing the software.

1.	Patient orientation (0020,0020) is a required (Type 2) field (part of the image module), but has little (if any) meaning for x-ray angiography.  More comprehensive anatomic and positioning information can be given for angiography.  The field cannot be omitted, as it is Type 2.  This seems like a harmless situation, as it can be included and just be left blank (as was done with the XA IOD s for the demonstration).  Consideration should be given to eliminating it from the XA IOD.

2.	Examining files from several vendors, there appears to be an inconsistency in the byte order of the File Meta Information version number.  According to Part 10, Table 7.1─1:

This is a two─byte field where each bit identifies a version of this File Meta Information header.  Version 1 is identified by bit 0 (low order bit of the second byte in the order of transmission) set to 1.  Version 2 is identified by bit 1 (of the second byte) set to 1...

Part 10 s Table C─1 lists (0002,0001) File Meta Information Version as a Value Representation (VR) of OB, consistent with the description above.  Some vendors appear to have incorrectly interpreted the use of the field to be US (Unsigned Short), resulting in the byte order getting reversed.  Version 1 would then incorrectly be 01 00 instead of the proper byte order of 00 01.

3.	The documentation for Part 10 does not specify a transfer syntax for DICOMDIR. This information is buried in the annex to Part 11, and was missed in the alpha release of this software.  The transfer syntax is actually specified to be Explicit Little Endian.  The software has now been changed to reflect this specification.

4.	Part 10 s section B.7 describes the Icon Image (0088,0200), defined as  a Sequence which contains a single Item encapsulating the Data Set made of the Data Elements of the Icon Image.   What is the transfer syntax for icons?  It is assumed that the icon image should have the same format as the DICOMDIR file in which it appears (an not the XA IOD file it refers to).

5.	Study ID is a Type 1 field in the DICOMDIR, but a Type 2 field in the XA IOD.  Probably should be changed to be Type 1 in the XA IOD.

6.	Part 11 Draft version 0.7 dated September 7, 1994 lists additional keys for the DICOMDIR (section A.1.3.3).  Institution Name (0008,0080) and Institution Address (0008,0081) are listed as Type 2 fields, but are Type 3 in the XA IOD.  Similarly, the Performing Physician Name (0008,1050) is Type 2 in DICOMDIR but Type 3 in the XA IOD.  All should probably be changed to be Type 2 in the XA IOD.

7.	The Application Profile for Basic Cardiac X-Ray Angiographic Studies on CD R Media described in Part 11 (APL-XA-BCCD) does not allow the use of the biplane image object, requiring instead that the biplane acquisitions be transferred as two single plane SOP instances.  The information in the DICOMDIR is inadequate to be able to link two such SOP instances back together.  That is, there is no way of knowing at the directory (and possibly even with information with the files themselves) that two multi-frame images were taken together.

8.	When a file set cannot be written entirely onto one CD R, it is divided into two (or more) independent DICOM file sets (each with their own DICOMDIR).  There is currently no mechanism to link these two discs electronically, that is, no way of indicating in machine-readable form that the first disk is actually  1 of 2 .

9.	In Annex A to Part 11 (X-Ray Angiographic Application Profile), Table 2 mentions a Detached Patient Management IOD.  While it is not clearly stated as such, this is considered an optional component, and is not included in the demonstration.

10.	I have found what appears to be a significant error in the documentation with regard to the description of  encapsulation of Encoded Pixel Data (Part 5, section A.4).  This section indicates that the data set will be encoded with Explicit VR Little Endian transfer syntax.  However, Figure A.4-1 incorrectly shows the Pixel Data Element Tag (7FE0,0010) encoded with Implicit VR.  Four bytes are missing from the first panel of this diagram (the VR of OB, and a 2-byte filler of 0's).  This resulted in incorrect coding of the pixel specification in the alpha release of the software.  This has now been corrected, and the problem with the documentation reported to Working Group 5.

11.	The way that lossless JPEG encoding is used for a multi-frame image is not fully specified.  A JPEG byte stream can have a single file header followed by one or more images (each preceded with its own image header).  When encoding a multi-frame image, each frame could be treated as its own JPEG  file  (i.e., with a file header and image header for each frame).  Or, there could be a single file header, and each frame could have just its image header.  At first glance, the latter would seem to be the most logical approach.  However, this limits the use of the offset table in the encoded pixel specification to achieve random access of any frame.  Instead, the first frame would have to be accessed to get the file header information, then the desired frame accessed, reading its image header.  Furthermore, since the Huffman table is included as part of the file header, this approach would result in using the same Huffman table for all frames (thereby reducing the compression ratios possible).  For the demonstration software, only the first approach has been taken.  Each frame of a multi-frame image is a complete JPEG  file , with its own file header and image header.  The JPEG compression utility that was used to create the encoded pixels analyzed each frame individually, and provides an optimized Huffman table for each frame.  Either the standard should be modified to clarify how the JPEG encoding should be done for multi-frame images, or File Set Readers should anticipate both possible scenarios.

12.	According to Part 5, Section A.2 the VR of the Icon Image Pixel Data (7FE0,0010) is OW when the Transfer Syntax is Little Endian Explicit VR (as in the DICOMDIR).  But, Part 10, Section B.7 restricts the icon pixel data to 1 or 8 bits allocated, and 1 or  8 bits stored.  Pixels would be a maximum of one byte (8 bits) long.  Some uncertainty remains in how to best store the pixels in this situation.  I have elected to follow the procedures outlined in the DICOM  Change Proposal CP─14 , that has a status of  interim approved .  This change affects Part 5, Section A.2 (among others), and states that data element (7FE0,0010) can have the Value Representation of OB when bits allocated (0028,0100) has a value less than or equal to 8.  This is also stated in the revised version of Part 5, Section 8, which says:

If Pixel Data is sent in a Native Format, the Value Representation OW is most often required.  The Value Representation OB may also be used for Pixel Data in cases where Bits Allocated has a value less than or equal to 8, but only with Transfer Syntaxes where the Value Representation is explicitly conveyed.

Implementors who are interested can get a copy of the complete Change Proposal from the ftp site ftp.xray.hmc.psu.edu (150.231.1.230) in the dicom_docs/drafts subdirectory.  The document is CP14_interm_approv.hqx.

13.	There still appears to be some confusion and disagreement concerning the status of the Change Proposal described above.  The latest word we have is that is has been accepted.  For the purposes of this demonstration, pixel data will be represented as OB when bits allocated (0028,0100) has a value less than or equal to 8 (for example, icon images).

14.	There is an inconsistency in the printed DICOM standard with regard to (0018,1063) Frame Time and (0018,1065) Frame Time Vector.  Part 3, sections C.7.6.5.1.1 and C.7.6.5.1.2 respectively refer to these as seconds.  These times are actually specified in milliseconds (as correctly specified elsewhere in the document).  All times in the files used in this demonstration are in milliseconds (for example a value of 33 for a 30 frame-per-second cine run).

15.	Icon images are required in the DICOMDIR for XA images, but not for Ultrasound (US) images.  Since US images will be included on the CD R for the demonstration, icon images have been generated and are included in the corresponding DICOMDIR entries for the US images.  Since XA icons are 128 by 128, the US icons that were generated are the same size.  More typically, US icons (as with other image icons in a DICOMDIR) would be 64 by 64.  According to Part 10, Section B.7,  Every DICOM Compliant [File Set Reader]... shall either ignore Icon Images or support them up to a pixel matrix of [64 by 64].  Icons larger than 64 by 64 may be stored by [a File Set Creator] but may be ignored by [File Set Readers].

15.	Ultrasound IOD s now use Large Palette data (16-bits per entry) for PALETTE COLOR images.  There is no accepted standard for how to translate a 24-bit RGB image to an 8-bit PALETTE COLOR image with corresponding Large Palette LUT s.  It is the intent of these 16-bit LUT's that the entire range of 0x0000 to 0xffff be used.  Some implementors have taken 24-bit RGB data (which has 8 bits of information for each color, not 16), and assigned entries in the 16-bit LUT's by padding the value on the left with a zero (that is 0xff becomes 0x00ff).  Other implementors have chosen padding on the right (0xff becomes 0xff00).  In keeping with the spirit of using the entire range of values, we have chosen to  byte replicate  when creating the LUT values.  For example, 0xff becomes 0xffff; 0x23 becomes 0x2323, etc.  Implementors reading these three 16-bit LUT s for display on 24-bit hardware (or software) can therefore take either byte of the 16-bit value.  The most general approach of all for a reader is to examine all 256 entries in the table, find the lower value and the highest value present, and compute a scaling factor appropriate for the target hardware.

Frequently Asked Questions (FAQ).c.Frequently Asked Questions (FAQ);

It is a common practice in the world of the InterNet to put together a list of Frequently Asked Questions (referred to simply as an  FAQ ) about a particular topic, along of course, with the answers.  This seemed like a convenient way to clarify some of the issues surrounding the DISC 95 DICOM demo.  Okay, so maybe some of these questions haven t been asked frequently (or maybe not at all), but they might be.  Anyway, here is my attempt at a DISC 95 FAQ:

1.	Software

1.1	Is the software in the public domain?

The meaning of  public domain  is a common source of misunderstanding in the computing community, especially as bulletin boards and networks become more common.  The software is copyrighted by the Brown University Institute for Medical Computing, and the American College of Cardiology.  Derivative works (especially those using specific portions of code) are expected to include the appropriate copyright notices (should appear in the source code, be compiled into the binaries, and should display on the computer screen as the program is started).  This copyright process is different from code that is released with no restrictions.

Vendors who are participating in the demonstration at DISC 95 will have a considerable lead time in that the software is provided privately to these vendors starting in October, 1994.  The release to the general imaging community will not occur until well after the DISC 95 demo.  At that time, it will be made widely available through InterNet distribution.

1.2	Has this software been validated?

Against what?  There is little work that has been done (at least that has been publicly acknowledged) in the area of XA IOD files or in creating DICOMDIR specifications.  The staged release program carried out by the ACC is the validation process.  What better way to validate the data and the programs than to have over 25 vendors dissecting, testing, verifying, examining, and commenting!  Those vendors with existing programs that have XA IOD capability or DICOMDIR capability should be comparing their findings and results with those provided with this distribution.  The content of the DICOM files and the behaviors of the programs have been reviewed and examined extensively.

1.3	How can I contact the Brown University Institute for Medical Computing with feedback or questions?

Electronic mail works very well (Jonathan_Elion@brown.edu), and FAX also works well (401-331-8501).  These are checked frequently, and turn-around time on questions can be expected to be short.

1.4	Why isn t there a complete working demo provided (with display capabilities)?

What has been provided is a  shell  or  stub  of a demo program, mainly to show how the files are manipulated, and how information from the DICOMDIR directory is handled.  For the DISC  95 demonstration, it was felt that it was important that all vendors start from a  level playing field .  It was not possible to anticipate all possible hardware and software environments.  If one platform had been selected for a complete working version of the demo (for example, X-Windows), than that would have given an unfair advantage to any vendor developing for that environment, while not helping vendors developing for Windows (for example).

1.5	What platforms are supported  by the demo software?

The programs have been developed under DOS 6.2 using the Borland C compiler, version 4.0.  Every effort has been made to restrict the program code to standard ANSI C, with no operating system, Posix, or other language or system quirks.  Extensive use of type casting, together with strict error and syntax checking in the Borland environment has helped to minimize warnings on other systems (but may not totally eliminate compiler warnings). The program senses whether it is running on a Little-Endian (Intel-like) or Big-Endian (Motorola-like), and handles the data files accordingly (it is not necessary to modify the program or any #define statements to get it to compile on  the different Endian systems.  The programs have also been tested under Unix on a Sun workstation (SunOS 4.1.3 on a Sparc 10), and under Macintosh System 7.1 with CodeWarrior 4.5 (with ANSI C library).  Just be aware that what one compilers considers to be  ANSI C  may not be consistent with the interpretation of another compiler.

2.	Image Files

2.1	Why bother distributing the build program?  Isn t it enough to have the XA IOD files?

Good question (glad I thought of it!).  I have always felt that no matter how complete documentation is,  it is always helpful to see the program that built a file.  If there are any questions about the contents of an XA IOD file, you can examine the contents of the file itself (using the parse utility), the data file that created it (XA000001.SCP, for example), or the program that built it (build.c).  These are felt to be valuable supplements to the documentation, and particularly helpful to those just beginning their DICOM development efforts.

2.2	How were the images selected for the demo disc?

Images were submitted by vendors, hospitals, and other investigators.  ACC members reviewed these images for image quality and clinical interest, and selected a representative assortment from this list.

2.3	What types of ultrasound images are on the demo disc?

Unlike X-Ray Angiography that currently only defines one type of image, the ultrasound part of the DICOM standard allows for several different formats.  For the sake of simplicity, only a few of these formats are represented on the demo disc (none use compression).  These include MONOCHROME2 (8-bit monochrome images), PALETTE COLOR (8-bit  indexed color  with corresponding color lookup tables), and RGB (24-bit pixels, with red, green, and blue triplets for each pixel).

3.	CD R and CD ROM Issues

3.1	I thought that the standard interchange medium was to be CD R.  Why then did the ACC use CD ROM for the DISC 95 demo?

As explained at the beginning of this documentation, the standard interchange medium is CD R, meaning that a single platter is recorded by a single writing device.  The disk that is produced is compatible by standard CD ROM readers, and the same data can also be transcribed using the manufacturing techniques developed for traditional CD ROM s.  This demonstration used a large-scale  pressing  of 5,000 CD ROM s of the file set so that adequate numbers of the discs are available for meeting attendees, vendors, and ACC officials.  The differences between a CD R and a CD ROM are technical with regards to their method of recording, and will not effect the reading of the discs.

3.2	How does the ISO 9660 file system play into this scheme?

Your operating system should be capable of handling the ISO 9660 file system (drivers are available for DOS, Windows, Unix, and Mac environments).  This means that the files on the CD R can be read by the same software routines used during testing on disk-based files.  There are some quirks of file name conversions unique to each operating system and device driver.  For example, Unix tends to convert upper case ISO 9660 file names to all lower case, and preserve the final  .  of the file name.  DOS and Windows don t care about upper or lower case, or about the final  .  of the file name.  Macintosh seems to preserve uppercase, and strips out the final  .  of the file names.

4.	JPEG Compression

4.1	What support is there for lossless JPEG compression and decompression?

There is excellent software available in the public domain to handle software compression and decompression.  Some vendors already have this software capability, and some have the hardware capability.  The public domain software will fill the gap for now for vendors without this capability.

4.2	What about lossy compression?

Some vendors have started to show systems that use lossy JPEG compression to achieve higher compression ratios, more compact file storage, and higher speed access of files on compact discs.  This is clearly outside of what has been specified by the DICOM standard.  Lossless JPEG compression is the only format currently specified for X-Ray Angiographic images in the DICOM standard.  Further scientific studies are needed to assess the impact of lossy compression on the diagnostic content and clinical utility of angiographic images.